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AirCurve 11 VAuto
Care, Maintenance & Staff Sign-Off
Daily • Weekly • Regular Review • 6-Month Filter Replacement
	Person receiving care:
	____________________________
	Date of birth:
	________________

	Staff member:
	____________________________
	Date:
	________________

	Shift / time:
	____________________________
	Location / room:
	________________


DAILY / BEFORE USE — ResMed care instructions + operational check
	Check
	Action required
	✓ / ✗
	Initials / comments

	1. Humidifier tub
	Empty any remaining water. Wipe thoroughly with a clean disposable cloth and allow to dry out of direct sunlight.
	☐ Pass
☐ Fail
	

	2. Refill water tub
	Refill before use with the correct water for the fitted tub. HumidAir 11 Standard: distilled water only. Do not exceed the MAX line.
	☐ Pass
☐ Fail
	

	3. Visual condition
	Check device, tub, tubing, mask and connections for obvious damage, contamination or deterioration.
	☐ Pass
☐ Fail
	

	4. Setup
	Ensure tub and tubing are correctly seated/connected and the mask vent is unobstructed.
	☐ Pass
☐ Fail
	

	5. Power / start-up
	Confirm the device powers on normally and there is no unresolved warning or fault.
	☐ Pass
☐ Fail
	

	6. Prescribed therapy
	Confirm the prescribed therapy is available. Do not alter prescribed clinical settings unless authorised.
	☐ Pass
☐ Fail
	


WEEKLY — ResMed cleaning schedule
	Check
	Action required
	✓ / ✗
	Initials / comments

	1. Air tubing
	Wash in warm water using mild dishwashing liquid. Rinse thoroughly and allow to dry out of direct sunlight or heat.
	☐ Pass
☐ Fail
	

	2. Humidifier tub
	Wash in warm water with mild dishwashing liquid, OR use 1 part household vinegar to 9 parts water at room temperature. Rinse thoroughly and air-dry.
	☐ Pass
☐ Fail
	

	3. Outlet connector
	Wash in warm water with mild dishwashing liquid, OR 1 part household vinegar to 9 parts water. Rinse thoroughly and air-dry.
	☐ Pass
☐ Fail
	

	4. Device exterior
	Wipe the exterior of the device with a dry cloth.
	☐ Pass
☐ Fail
	

	5. Mask / headgear
	Clean and inspect according to the specific mask manufacturer's user guide.
	☐ Pass
☐ Fail
	


REGULAR INSPECTION — ResMed recommends regular checks
	Check
	Action required
	✓ / ✗
	Initials / comments

	1. Humidifier tub
	Replace if leaking, cracked, cloudy or pitted, or if the seal is cracked/torn. Remove white deposits with 1 part household vinegar to 9 parts water, then rinse.
	☐ Pass
☐ Fail
	

	2. Air tubing
	Replace if there are holes, tears or cracks.
	☐ Pass
☐ Fail
	

	3. Air filter
	Inspect for holes, dirt or dust blockage. The filter is not washable or reusable.
	☐ Pass
☐ Fail
	

	4. Device performance
	Stop use and escalate if the device is damaged, does not perform as usual, or makes unusual sounds.
	☐ Pass
☐ Fail
	

	5. Bacterial/viral filter, if fitted
	Regularly check for moisture or contamination; follow local policy and manufacturer instructions.
	☐ Pass
☐ Fail
	


MONTHLY FACILITY REVIEW — local governance check (not a separate ResMed monthly cleaning requirement)
	Check
	Action required
	✓ / ✗
	Initials / comments

	1. Records
	Confirm daily/weekly checks are being completed and documented.
	☐ Pass
☐ Fail
	

	2. Supplies
	Confirm spare filter and required cleaning/replacement supplies are available.
	☐ Pass
☐ Fail
	

	3. Prescription / care plan
	Confirm current prescribed therapy and escalation plan are available to staff.
	☐ Pass
☐ Fail
	

	4. Equipment review
	Review recurring leaks, alarms, skin issues, condensation, damage or tolerance problems and escalate as required.
	☐ Pass
☐ Fail
	


EVERY 6 MONTHS — ResMed air-filter replacement
	Check
	Action required
	✓ / ✗
	Initials / comments

	1. Replace air filter
	Replace the AirCurve 11 air filter every 6 months, or sooner if it has holes or is blocked by dirt/dust. Do not wash or reuse it.
	☐ Pass
☐ Fail
	


FAULT / ESCALATION RECORD
	Fault / concern:
	____________________________________________________________

	Action taken / escalated to:
	____________________________________________________________

	Staff signature / time:
	____________________________________________________________


Training/staff aid: use with the individual care plan, local infection-prevention policy, the mask instructions and the current ResMed AirCurve 11 user/clinical guidance. For multi-patient reuse, follow the applicable ResMed reprocessing instructions.
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