AirCurve 11 VAuto
Care, Maintenance & Staff Sign-Off

Daily • Weekly • Regular Review • 6-Month Filter Replacement

| Person receiving care: | ____________________________ | Date of birth: | ________________ |
| --- | --- | --- | --- |
| Staff member: | ____________________________ | Date: | ________________ |
| Shift / time: | ____________________________ | Location / room: | ________________ |

DAILY / BEFORE USE — ResMed care instructions + operational check

| Check | Action required | ✓ / ✗ | Initials / comments |
| --- | --- | --- | --- |
| 1. Humidifier tub | Empty any remaining water. Wipe thoroughly with a clean disposable cloth and allow to dry out of direct sunlight. | ☐ Pass ☐ Fail |  |
| 2. Refill water tub | Refill before use with the correct water for the fitted tub. HumidAir 11 Standard: distilled water only. Do not exceed the MAX line. | ☐ Pass ☐ Fail |  |
| 3. Visual condition | Check device, tub, tubing, mask and connections for obvious damage, contamination or deterioration. | ☐ Pass ☐ Fail |  |
| 4. Setup | Ensure tub and tubing are correctly seated/connected and the mask vent is unobstructed. | ☐ Pass ☐ Fail |  |
| 5. Power / start-up | Confirm the device powers on normally and there is no unresolved warning or fault. | ☐ Pass ☐ Fail |  |
| 6. Prescribed therapy | Confirm the prescribed therapy is available. Do not alter prescribed clinical settings unless authorised. | ☐ Pass ☐ Fail |  |

WEEKLY — ResMed cleaning schedule

| Check | Action required | ✓ / ✗ | Initials / comments |
| --- | --- | --- | --- |
| 1. Air tubing | Wash in warm water using mild dishwashing liquid. Rinse thoroughly and allow to dry out of direct sunlight or heat. | ☐ Pass ☐ Fail |  |
| 2. Humidifier tub | Wash in warm water with mild dishwashing liquid, OR use 1 part household vinegar to 9 parts water at room temperature. Rinse thoroughly and air-dry. | ☐ Pass ☐ Fail |  |
| 3. Outlet connector | Wash in warm water with mild dishwashing liquid, OR 1 part household vinegar to 9 parts water. Rinse thoroughly and air-dry. | ☐ Pass ☐ Fail |  |
| 4. Device exterior | Wipe the exterior of the device with a dry cloth. | ☐ Pass ☐ Fail |  |
| 5. Mask / headgear | Clean and inspect according to the specific mask manufacturer's user guide. | ☐ Pass ☐ Fail |  |

REGULAR INSPECTION — ResMed recommends regular checks

| Check | Action required | ✓ / ✗ | Initials / comments |
| --- | --- | --- | --- |
| 1. Humidifier tub | Replace if leaking, cracked, cloudy or pitted, or if the seal is cracked/torn. Remove white deposits with 1 part household vinegar to 9 parts water, then rinse. | ☐ Pass ☐ Fail |  |
| 2. Air tubing | Replace if there are holes, tears or cracks. | ☐ Pass ☐ Fail |  |
| 3. Air filter | Inspect for holes, dirt or dust blockage. The filter is not washable or reusable. | ☐ Pass ☐ Fail |  |
| 4. Device performance | Stop use and escalate if the device is damaged, does not perform as usual, or makes unusual sounds. | ☐ Pass ☐ Fail |  |
| 5. Bacterial/viral filter, if fitted | Regularly check for moisture or contamination; follow local policy and manufacturer instructions. | ☐ Pass ☐ Fail |  |

MONTHLY FACILITY REVIEW — local governance check (not a separate ResMed monthly cleaning requirement)

| Check | Action required | ✓ / ✗ | Initials / comments |
| --- | --- | --- | --- |
| 1. Records | Confirm daily/weekly checks are being completed and documented. | ☐ Pass ☐ Fail |  |
| 2. Supplies | Confirm spare filter and required cleaning/replacement supplies are available. | ☐ Pass ☐ Fail |  |
| 3. Prescription / care plan | Confirm current prescribed therapy and escalation plan are available to staff. | ☐ Pass ☐ Fail |  |
| 4. Equipment review | Review recurring leaks, alarms, skin issues, condensation, damage or tolerance problems and escalate as required. | ☐ Pass ☐ Fail |  |

EVERY 6 MONTHS — ResMed air-filter replacement

| Check | Action required | ✓ / ✗ | Initials / comments |
| --- | --- | --- | --- |
| 1. Replace air filter | Replace the AirCurve 11 air filter every 6 months, or sooner if it has holes or is blocked by dirt/dust. Do not wash or reuse it. | ☐ Pass ☐ Fail |  |

FAULT / ESCALATION RECORD

| Fault / concern: | ____________________________________________________________ |
| --- | --- |
| Action taken / escalated to: | ____________________________________________________________ |
| Staff signature / time: | ____________________________________________________________ |

Training/staff aid: use with the individual care plan, local infection-prevention policy, the mask instructions and the current ResMed AirCurve 11 user/clinical guidance. For multi-patient reuse, follow the applicable ResMed reprocessing instructions.